## Slide 1
Flood Victim Triage System
Low-cost vital-sign screening with ESP32 and MAX30102 for disaster
response
Red, yellow and green priority indicators

## Slide 2
Why this project?
1
The problem
Few doctors, many victims.
Without triage, critical
patients wait in the same
queue as minor cases.
2
My idea
Volunteers screen each
victim in about a minute
using a low-cost sensor and
a simple web page.
3
The result
Red cases are flagged for
immediate care. Records
are stored offline and sent
to doctors when a network
returns.
2

## Slide 3
System architecture
MAX30102
sensor
Reads SpO₂ and
heart rate
ESP32
microcontroller
Wi-Fi access point
and web server
Volunteer
webpage
Form, triage colour,
GPS
Device
storage
Saved offline in the
browser
Medical
server
Receives records
when online
Entered by the volunteer: temperature, blood pressure, consciousness and notes
Fed back to the device: the priority colour lights the red, yellow or green LED
3

## Slide 4
Electrical components used
ESP32 dev board
Dual-core microcontroller,
3.3 V logic, built-in Wi-Fi
MAX30102 sensor
Red and infrared LEDs with a
photodetector, I2C interface
Status LEDs
Red, yellow and green
priority indicators
220 Ω resistors
Limit each LED current to
about 6 mA
Breadboard and
jumpers
Solder-free wiring for quick
prototyping
Power supply
USB power bank or 5 V
charger for field use
4

## Slide 5
Working principle
SpO₂ and heart rate by pulse oximetry
•Oxygen-rich blood absorbs more infrared light;
oxygen-poor blood absorbs more red light
•The sensor measures reflected red and infrared light
from the fingertip
•SpO₂ comes from the ratio of the pulsing to steady
signal at each colour
•Heart rate comes from the time between pulse peaks
LED current check
≈ 6 mA
I = (3.3 V − 2.0 V) ÷ 220 Ω
Safe for an ESP32 pin and bright
enough for an indicator
5

## Slide 6
Frontend, backend and database
Frontend
HTML, CSS and JavaScript
page: patient form, live
sensor readout, triage
colour, GPS capture and
patient list.
Backend
ESP32 web server provides
/vitals and /led. An optional
Node.js server receives
records at /api/patients.
Database
Browser localStorage keeps
records offline. SQLite on
the medical server stores
synced records.
Data format: JSON Link: Wi-Fi or USB Sync: HTTP POST when online
6

## Slide 7
How priority is decided
RED
Immediate: SpO₂ below 90%, heart rate above 130 or below 40, systolic BP
below 90, temperature 104 °F or above or below 95 °F, or unresponsive
YELLOW
Urgent: SpO₂ 90 to 93%, heart rate 111 to 130 or 40 to 49, systolic BP 90 to 99
or above 180, temperature 101.3 to 103.9 °F or 95 to 96.7 °F, or reduced
consciousness
GREEN
Stable: all readings in the normal range and the patient is alert
Rule: the worst single reading decides the colour. Thresholds are starting values and must be validated by medical
staff.
7

## Slide 8
Programming languages and tools
C / C++
Firmware in Arduino IDE: sensor reading, SpO₂ and heart-rate algorithm, Wi-Fi
server, LED control
HTML and CSS
Page structure and a layout that works on phone and laptop
JavaScript
Triage rules, ESP32 link (fetch and Web Serial), GPS, offline storage, sync
JSON
Simple data format between ESP32, webpage and server
Node.js and SQL
Optional medical-department server and database
8

## Slide 9
Advantages and limitations
Advantages
•Low cost, built from common parts
•Works offline and syncs later
•Fast screening, about one minute per
patient
•Automatic GPS location for every victim
•LED colours readable without the screen
Limitations
•MAX30102 is not medical grade; motion
and cold fingers affect readings
•Temperature and blood pressure are still
manual
•Records can be lost if browser data is
cleared
•No login or encryption yet
•Triage thresholds need clinical validation
9

## Slide 10
Conclusion and future scope
A low-cost, offline-first tool that
helps volunteers find the
victims who need a doctor first.
Next steps
•Add a temperature sensor such as MLX90614
•Battery pack and waterproof enclosure
•Login and encrypted sync
•Doctor dashboard with a map
•Clinical validation of thresholds

## Slide 11
“We cannot stop natural disasters, but we can arm ourselves with knowledge: so many lives wouldn't be lost if there was enough disaster preparedness.” — Petra Nemcova

## Slide 12
Thank You